M.P. State Cooperative Dairy Federation Ltd.

Dugdha Bhavan, Habibganj, Bhopal - 462 024 (M.P)

Telephone No. : 0755 2602145 ;   Fax No. 0755 2583149

e-mail : ho.mpcdf@nic.in  website : www.mpcdf.gov.in

Notice Inviting e-Tender for Supply of

Lab Equipments for testing of Milk and Milk Products

Tender Reference No. MPCDF/PUR-288/2020/OT-1

Bid submission starts from :   Date  12.06.2020 on             MP e-portal  www.mptenders.gov.in

Cost of Tender Form: Rs. 1000/-

M.P. State Cooperative Dairy Federation Ltd.

Dugdha Bhavan, Habibganj, Bhopal - 462 024 (M.P)

Telephone No. : 0755 2602145 ;   Fax No. 0755 2583149

e-mail : ho.mpcdf@nic.in  website : www.mpcdf.gov.in

E-Tender Notice

MP State Cooperative Dairy Federation Ltd. (MPCDF), Dugdha Bhavan, Habibganj, Bhopal, an apex cooperative organization in Madhya Pradesh, invites tender from reputed manufacturers / suppliers for supply of Quality Control Lab Equipments, kit and strips for testing of Milk and Milk Products strictly adhering to the terms and conditions prescribed in the tender document.

The cost of tender document is Rs.1000/-(Rs.One Thousand only).

The tender document can be obtained from the site of MP e-portal    www.mptenders.gov.in on remitting online payment of the cost. The details relating to the tender can also be referred at MPCDF website www.mpcdf.gov.in.

Key Information

| Lab equipments | EMD required Rs as per Lab equipments | Last date and time of purchase of tender form | Last date and time of bid submission | Date ,time & place of opening technical bid | Date of opening of Financial bid |
| --- | --- | --- | --- | --- | --- |
| Automatic Kjeldahl Unit | 80000 /- | 3.06.2020 2:30 PM | 3.06.2020 3:00 PM | 04.06.2020 03:00 PM at MPCDF Bhopal | 06.06.2020 03:00 PM at MPCDF Bhopal |
| Sodium Potassium Analyzer | 50000/- |  |  |  |  |
| Melamine strip | 25000/- |  |  |  |  |
| Adultration kit | 1000/- |  |  |  |  |

Managing Director

M.P. State Cooperative Dairy Federation Ltd.

Dugdha Bhavan, Habibganj, Bhopal - 462 024 (M.P)

Telephone No. : 0755 2602145 ;   Fax No. : 0755 2583149

e-mail : ho.mpcdf@nic.in  website : www.mpcdf.gov.in

Tender Document

| Schedule I | General Terms and Conditions |
| --- | --- |
| Schedule II | Bidder's details |
| Schedule III | Specification of the lab equipments |
| Schedule IV | Price Bid |
| Last date and time of purchase of tender form | 03.06.2020 up to 2:30 PM |
| Last date for online upload of Tender Form | 03.06.2020 up to 3.00 PM |
| Venue for submission of Tender documents : |  |
| (i) Technical Bid & Price bid online (mandatory) | (i) Online at www.mptenders.gov.in by 03.06.2020 up to 3.00 PM |
| Date & Time of Opening Technical Bid and EMD | 04.06.2020 at 03.00 PM |
| Date & Time of Opening Financial Bid | 06.06.2020 at 3:00 PM |
| Venue of bids opening | O/o M P State Cooperative Dairy Federation Limited, Dugdha Bhawan, Habibganj, BHOPAL - 462024 (MP) |
| Cost of Tender Form | Rs. 1000/- ( Rupees One thousand only ) |

Managing Director

SCHEDULE - I

General Terms & Conditions of Tender

###### MP State Cooperative Dairy Federation Limited (MPCDF), Dugdha Bhavan, Habibganj, Bhopal (MP) invites online tenders through MP e-tender portal www.mptenders.gov.in from bonafide manufacturers / suppliers for supply of Quality control lab equipment for testing of milk and milk products for  dairy plants, mini dairy plants and chilling centers which comes under the milk Unions of Bhopal, Indore, Gwalior  and Jabalpur strictly adhering to the quality benchmarks and technical specifications prescribed in the tender document. As per the list of items/equipments Automatic Kjeldahl Unit, Sodium Potassium Analyzer, Melamine Strips and Adulteration Kit supplier or manufacturer can participate in tender for supply of any of these items he manufacturer/supply or all of them. Technical validation of each of the items will be done and thereafter the financial bid will be open as per the MP e tenders guidelines.

###### The Managing Director, MP State Cooperative Dairy Federation Ltd. Bhopal solely reserves the rights to accept or reject, partially or fully, any or all tenders without assigning any reasons thereof.

1.0      DECLARATION :

  - Mere submission of a bid by a tenderer itself implies that he/she has read all the instructions, terms and conditions stipulated in the tender and has completely made himself / herself aware of scope, quality & technical specifications of lab equipments, supplies to be made at the defined destinations and, have fully satisfied himself / herself with all related bindings that were expected from him / her through conscious participation.

2.0      EMD & RATE VALIDITY PERIOD :

2.1   	The tenderer would be required to deposit prescribed Earnest Money. EMD submitted on line only. Without EMD any bidder shall not be entertained.

2.2	Price bid if approved shall be valid for a period one year from the date of rate approval.

3.0  	TENDER SUBMISSION :

3.1  	A tenderer will not submit more than one tender. No Individual or firm shall be allowed to submit more than one tender under different names.

3.2 	The tenderer will ensure to clearly indicate his / her position / capacity in the bidding firm in signing the tender and related documents.

3.3	The Bidder should posses a valid GST registration certificate and PAN of income tax of previous two years.

3.4	The tenderer have to submit Earnest Money Deposit (EMD) online only. In no case the other modes of EMD deposition would be entertained if submitted physically by any bidder.

3.5 	The tenderer will have to submit the technical bid papers online only.

3.6	The Financial bid shall mandatorily be submitted online only (refer Schedule III). In no case the financial bids would be entertained if submitted physically by any bidder.

3.7	The EMD will be opened online. Once it is found in order then only then the  envelope containing the technical bid and scanned EMD copy of the concerning bidder shall be opened for further scrutiny.

3.8 	The technical bid shall be opened as per the mentioned schedule. Bidders may participate in the technical bid physically or via video conferencing by informing the coordinator of the tender by providing their e mail and mobile numbers.

3.9	The Financial bid of a bidder will be opened only when the individual tenderer / firm has successfully qualified the technical bid parameters / criteria.

4.0	OTHER PRE-REQUISITES :

4.1 	The tenderer should clearly indicate their mailing address, telephone, mobile, fax, e-mail, PAN and GST number. Any change in the address should immediately be communicated to the Managing Director MP State Cooperative Dairy Federation Ltd. Bhopal for exchanging future correspondence.

4.2	Negligence or carelessness on the part of the tenderer in filling the tender form / price / quote etc. shall not lend him / her any rights to withdraw from the tender once it is opened.

4.3	The tenderer shall submit the tender in the prescribed form only. Conditional tenders, in any case, are liable for outright rejection.

4.4	The tenderer will ensure to sign at the bottom of each page of the tender document in token of their acceptance to the terms and conditions of the tender.These scanned copies of the signed document must be up loaded.

4.5	The tenderer cannot sublet or assign wholly or partially the approved supply order to any other party/parties. In any case, if it is found so, then the assigned supply order will immediately be cancelled and the party will be solely liable for black listing by MPCDF.

4.6	A copy of manufacturing license is to be attached with other technical bid documents of tender for the equipments Automatic Kjeldahl Unit and Sodium Potassium Analyzer.  For strips and kit if  bidder is not manufacturer then the authorized dealers / channel partner etc. dealing with the item will have to attach a copy of authorization certificate of the concerning manufacturer.

4.7	The tenderer will have to  provide all the details of his/her unit in the forms given at Schedule-III.

4.8   Bidders must upload previous 2 years experience by giving the proof of equipments, kit and strip they have supplied. Purchase order copy and satisfactory certificate of customers must be up loaded.

5.0 	EARNEST MONEY DEPOSIT (EMD) :

5.1 	Any tender not accompanied with EMD shall be liable for rejection. Earnest Money Deposit of unsuccessful tenderers will be returned from the e procurement side automatically as per the MP e tender norms.

5.2	The EMD of the successful tenderer will be released only on completion of supplies, testing and warranty thereby covering replacement and after one year of rate validity period.

5.3	No interest will be paid on the Earnest Money for the period during which (the EMD) lies in deposit with the Managing Director, MP State Cooperative Dairy Federation Ltd.  Bhopal.

6.0 	PRICES :

6.1	  The tenderer is required to quote price of equipment  as per below scope of work with all consumable parts on FOR destination of Milk Unions located Bhopal / Indore / Gwalior/Jabalpur and heir Mini Dairy plants and Chilling centers as shown below in table 1.1

| Sno | Lab equipments | Qty Required | Scope of Work |
| --- | --- | --- | --- |
| 1. | Automatic Kjeldahl Unit | 11 | Supply at FOR as per table 1.1, installation, commissioning , with 3 year AMC after warranty of min 1 year training to the lab personnel ,Consumable part as described in Schedule III. |
| 2. | Sodium Potassium Analyzer | 11 | Supply at FOR as ,er table 1.1, installation, commissioning , , with 3 year AMC after warranty of min 1 year,training to the lab personnel, Consumable part as described in Schedule III |
| 3. | Melamine strip | 11 | Supply at FOR as per table 1.1, training to the lab personnel |
| 4 | Adulteration kit | 11 | Supply at FOR as per table 1.1, training to the lab personnel. |

FOR Places for Lab Equipments  Table1.1

| SNo. | Name | FOR Place under MPCDF | Quantity required |
| --- | --- | --- | --- |
| 1 | Indore Milk Union | 1. Indore Main Dairy Plant – Indore. | 1 |
|  |  | 2.Khargone Mini Dairy Plant | 1 |
|  |  | 3.Jhabua Mini Dairy Plant. | 1 |
|  |  | 4.Khandwa Mini dairy Plant | 1 |
|  |  | 5.Burhanpur Chilling Center | 1 |
|  |  | 6.Sendhwa Main dairy Plant. | 1 |
|  | Bhopal Milk Union | - Betul mini dairy Plant | 1 |
|  | Gwalio Milk union | 1.Shivpuri Mini dairy plant | 1 |
|  | Jabalpur Milk union | 1.Chindwada Mini Dairy plant. | 1 |
|  |  | 2.Rewa Mini Dairy Plant. | 1 |
|  |  | 3.Singroli Chilling Center | 1 |
|  | Total |  | 11 |

7.2	Correction of errors, if any, in the tender document should be duly followed with initials/signature of the tenderer / firm at the places of every such correction made.

7.3	The prices offered by the tenderer should be firm and free from all anticipated escalations. In the dare validity period of one year for the date of tendering of opening of tender.

7.4	The lowest rate shall not be the only criteria in particular for awarding the tender until the supplier satisfies the stakeholder specifications and requirements.

7.5    The lowest rate so finalized will be inclusive of equipment cost at FOR basis and AMC cost for 3 years after warranty period of equipments.

7.6    After successful opening the tender, rates will be finalized and approved rates will be given by MPCDF within 30 days and the Purchase process i.e.  Signing of MOU, issuing of the Purchase order, payment and receipt of the Materials will be done by the Bhopal, Indore , Gwalior and Jabalpur Milk Unions.

8.0 	MODE OF DISPATCH:

8.1 	The affiliated milk union will place purchase order as per requirement during the tender validity period on the tenderer whose rate has been approved.

8.2   	The equipments to be supplied by successful tenderer shall be dispatched to the dairy plants under intimation to the purchaser with proper packaging / crating to avoid damages during transit.

8.3 	Wharfage / demurrage etc. on account of incorrect or delayed dispatch of material of documents shall be the responsibility of supplier and shall be recovered from the bill. The supplier shall bear the insurance charges of the equipments during transit.

8.4	Delivery schedule : The equipment  should be supplied in the stipulated time as per the below table 8.4.1 from the date of purchase order of the respective milk union.

Table 8.4.1

| No. | Equipments | Delivery schedule |
| --- | --- | --- |
| 1 | Automatic Kjeldahl Unit | 45 days |
| 2 | Sodium Potassium Analyzer | 45 days |
| 3 | Melamine Testing Strip | 30 days |
|  | Milk Adulteration Testing Kit | 30 days |

9.0 	LIQUIDATED DAMAGES :

9.1 	If for any reason the successful tenderer refuses to accept purchase order / fails to supply as per purchase order / backs out at a later date then the respective milk union(s) reserve the right to cancel the order and make purchases from other available alternate sources, at the risk and cost & expenses of the defaulting tenderer supplier and supplying firm could be blacklisted for all future dealings. Also the EMD/Security deposit would be forfeited by MPCDF

9.2   Once firm delivery schedule is fixed then it shall be strictly adhered to by the supplier. 	In case of supplies and services are not affected as per the schedules, the liquidated damages may be charged on the goods/services not so delivered as under.

| No. | Delay | Penalty |
| --- | --- | --- |
| 1 | Up to 15 days | 1% the Purchase order value. |
| 2 | Between 16 to 30 days | 2% Purchase order value |
| 3 | Beyond 30 days | Up to 5% the Purchase order value. |

.

9.3  	After supply of the equipment, the supplier should ensure that the test trials will be done  within 07 days of the and ensure starts smooth operation by guiding and providing complete training and details of the operation to the operator / concern person.

12.0   	PAYMENT :

12.1 	70% payment shall be released on safe receipt of equipment at site subject to acceptance after inspection by the concerned milk unions. The balance 20% payment shall be released only after completion and successful trial run of equipments and rest 10% payment will be released after submission of Bank Guarantee equal to 10% of order value towards performance guarantee of one year from the date of  installation  and commissioning.

13.0   AGREEMENT :

13.1	Signing of MOU by incorporating the required term and condition will be done between successful tender and the Chief Executive officers of Bhopal, Indore, Gwalior and Jabalpur Milk Unions.

13.2  	If the successful tendering firm / its authorized dealers commit breach of any of the conditions of tender even at a later date then the Managing Director, MP State Cooperative Dairy Federation Ltd., Bhopal reserves the right to cancel the award and authorize to purchase the material from other sources at the risk of the firm.

14.0  	DISPUTE ARBITRATION & FINAL AUTHORITY:

14.1 	It should be clearly understood that in the event of a successful tenderer failing to accept or / and execute the supply order, then decision of the Managing Director, MP state Cooperative Dairy Federation Ltd., Bhopal in this respect will be final and binding on the tenderer. For all disputes, the venue shall be at Bhopal, MP.

14.2	  Honorable  Pradhikrit adhikari/ Chairman MP State Federation Ltd Bhopal 																																																												will be the						Sole 				Arbitrator and his								decision 			shall  be							final and binding on 																																																														all 						the 										concerned.

14.3    For all disputes, the venue of the jurisdiction shall be at Bhopal, MP

Managing Director,

MP State Cooperative Dairy Federation Ltd.

Dugdha Bhawan,Habibganj,		               					       								Bhopal -462 024 (MP

Agreed to above terms & conditions of the tender

Signature & Seal of Tender

SCHEDULE – II

Technical Bid

#### : Form – A:

#### Details of the tenderer

I/We hereby furnish following particulars about our firm :

- Name of the Manufacturer/ supplier      ......................................................................

- Address of the Unit : 1. Office:  ..........................................................……………….

..........................................................................................................................

2. Factory: ....................................................................

..................................................................... ……………………...

- Names of Proprietor/Partner	  : 1.....................................................................................

.......................................................................................

- Name with designation of other person          :

authorized to sign the documents on              ......................................................................................

behalf of the unit,

(Authority letter should be produced)

- Telephone No/Mobile                                     	  :..................................................................................................

- Particulars of the Registration	  :No. ..................................... Dated........................................                    (attach photocopy certificate

issued by the competent authority)

- GST No. (attach photocopy)                          :..............................................................................................

- Permanent Account Number of

Income Tax (attach photocopy)	  :..................................................................................

- Whether the unit/firm or

sister concern or any	  :	               Yes/No (Please tick)

Unit of their Proprietorship, Partnership,

if any has been blacklisted/debarred or

penalized by any Central or

State Government/Cooperative/

Organization at any time.

- If the reply is ‘yes’ when & why give reasons in detail :……………………………………………….	  …………………………………………………………………………………………………………………………………………………………………………………………………………………………………………………..

- Please attach xeroxed  copy of last 2 years   :       .

of audited Income Tax Return

- EMD detail 	 :

1 Acknowledgement (Attached) ………………………………………... :...........................................................................................................................

2. Date …………….………….................................................................

3. Bank ...............................................................

4. Amount Rs. ...................................................\

13 Experience  of equipment supplied( not less than 2 year)

with satisfactory performance certificate from                                       `	clients..............................................................

I/We undertake that the information furnished in this tender documents is correct to the best of my/our knowledge and belief. Certified copies  of the above documents are enclosed.

Signature & Seal of the Tenderer

SCHEDULE – II

Technical Bid

#### : Form –B :

#### Details of the tenderer

I/we hear by furnish following particulars about our unit

  - If the unit earlier did any business                                  : Yes/No (Please tick)

in the past  with any of MPCDF affiliated milk unions

  - Are you supplying/supplied same equipments                    :……………………………………………

to some other State Cooperative                                            ………………………………………

Dairies?   If yes, to whom & when ?                                       ...........................................            Attach PO and Commissioning report

(Minimum 2 years)

  - List of current clients	:...........................................................

(Please attach a copy)			...........................................................

...........................................................

I/We undertake that the information furnished in this tender document is correct to the best of my/our     knowledge and belief. Certified copies of the above documents are  enclosed.

Signature & Seal of the Tenderer

SCHEDULE - III

Technical Specification of Lab Equipment

Tender must technically quote for the equipments he is biding covering all the technical details mentioned in Schedule III Form a, Form B , Form C and Form D.

| S. NO | Name of Equipments | Technical requirements of the lab equipments | Bidder must fill by "yes" or "No" bidding in below equipments. |
| --- | --- | --- | --- |
| I. | Automatic Kjeldahl Unit | SCHEDULE - III FORM-A |  |
| II. | Sodium Potassium Analyzer | SCHEDULE - III FORM -B |  |
| III. | Melamine Testing Strip | SCHEDULE - III FORM-C |  |
| IV. | Milk Adulteration Testing Kit | SCHEDULE - III FORM-D |  |

FORM-A

I.  Automatic Kjeldahl Unit

- Functional Requirement

- Estimation of protein content in milk and milk products. The offered equipment shall have   	digestion and distillation unit with automatic addition of reagents for efficient operation. 	Titration would be undertaken manually.

- Technical Specification

  - Digestion Unit

- Six place block digestion system able to hold 250 ml borosilicate glass tubes.

- Operational temperature: Ambient to 4500 C or higher.

- Display should show set value and current value of temperature and timer.

- It must have auto heating cut off on completion of digestion cycle.

- Efficient exhaust manifold system for removing acid fumes.

- Material of construction: Corrosion resistant

  - Scrubber unit

- Unit should have fume trap condensation unit followed by water and alkali neutralization.

- Alkali and water container capacity: 2 lit (Minimum)

- Operation pump must be made up of non-corrosive material.

- Unit shall have a mechanism to remove acid from exhaust air through dissolution in water   						    tank / carbon filter / any other suitable technique.

  - Distillation unit

- Unit shall have automatic steam injection.

- Unit shall have automatic water addition / dilution.

- Unit shall have automatic alkali and boric acid / acid addition.

- Unit shall give warning / indication for open door during operation and hold / deactivate 								the process.

- Unit shall have facility for adjustable steam power.

- Material of construction: Corrosion resistant.

- Chemical storage tanks shall be provided with minimum2.5 lit capacity.

- Auto cut off sensor for water circulation shall be provided

- Equipment shall be supplied with minimum 1-year warranty

- Unit shall be supplied with 12 borosilicate tubes (250 ml), required accessories, water connection pipes required for installation, demonstration and routine operations for at least 300 analyses.

6.0 	Demonstration:

i) Vendor must optimize digestion and distillation method for liquid milk and dried milk samples.

ii)Vendor must provide training to end user.

iii)Vendor must provide demonstration to end user including manual titration and verify the			distillation recovery of >99%.

7.0 	Others :

i) All the pre-requisite required for installation of the equipment shall be indicated by the vendor.

ii) AMC charges for the 2nd and 3rd year shall be indicated separately.

iii) It is the responsibility of vendor to supply a complete solution / unit for the analysis.

iv) Vendor should be able to supply all the spares / consumable required for repair at  least 8-10								years after installation.

FORM-B

II.  Sodium Potassium Analyzer :

| Equipment specifications |  |  |
| --- | --- | --- |
| S.No. | Parameter | Description |
| 1 | Unit of measurement | System shall display results directly in ppm. |
| 2 | Range of measurement | should measure Na, K in the range of : 2-200 ppm or better on both sides. |
| 3 | Sensitivity | Na: 2 ppm, K 1 ppm |
| 4 | Linearity | 3% or lower |
| 5 | Reproducibility | Less than 2 % Coefficient of Variation |
| 6 | Calibration | Minimum 5 point calibration facility should be present in the instrument.Relevant set of chemicals for calibration standards of Na (min. 100 ml of 100 ppm) and K (min. 100 ml of 100 ppm) should be provided by bidder. |
| 7 | Detector | Silicon photodiode or equivalent or better detection system |
| 8 | Auto flame ON/ OFF detection | Facility should be present in the system |
| 9 | Ignition System | Automatic ignition |
| 10 | Flame System | LPG/ PNG &Dry Oil Free Air |
| 11 | Gas Control | Provision for gas control should be available |
| 12 | Filter selection | Provision for filter selection should be available |
| 13 | Atomizer | Axial Flow type |
| 14 | Display | Alphanumeric LCD display or digital LED display |
| 15 | Power supply | Standard Indian - single phase |
| 16 | Air supply unit | The instrument should be supplied with air compressor unit to supply required gases. |
| 17 | Data storage and Printing | System shall be able to store data of at least 500 analysis. |
|  |  | RS 232 port for data transfer. All cables/ software for data transfer to be offered along with a suitable printer. |
| 18 | Accessories to be provided | Following accessories should be provided with the instrument: Two filters (Na and K) , calibration standards (as indicated at Sr. No. 6) , Atomizer Tube, Gas Lighter, Sample Beaker (Min.6 No), USB interface cable, LPG Tube, Air line PU Tube, power cord etc. All the other accessories required for smooth installation, demonstration and operation of the equipment for above a application should be provided. |
| 19 | Other | At the time of installation, the performance of the equipment should be demonstrated as per the provided specifications or OEM catalogue, whichever is better. |
|  |  | The accessories required for commissioning of the equipment shall be in the scope of supplier. |
| 20 | Training | Training to end user to be provided |
| 21 | Demonstration & Method Set up | Bidder has to standardize and optimize the for analysis of Sodium and Potassium at the installation site. |

FORM-C

III  Technical specifications for Melamine testing Strip :

- Kit must be working based on Immune-chromatographic technique.

- Limit of Detection of the Kit shall be within 80 to 120 µg/lit.

- Kit must be able to detect melamine in raw and processed milk.

- Vendor shall provide 650- 700 strips.

- Visual interpretation of the test results shall be possible with appearance of test and control 		lines.

- Vendor shall include additional requirements like incubator, auto-pipettes, tips or any other 		that is required for achieving accurate and reproducible results.

- Vendor shall provide positive and negative standard to the working of kit.

- Vendor shall provide staggered delivery in two to three shipment on mutually agreeable schedule in a period of 2 years, due to shelf life of strips.

- Vendor has to conduct verification of the kit to end user before or after the purchase process.

- Vendor shall provide demonstration and training to end user.

FORM-D

IV.  Specification of Milk Adulteration Testing Kit :

| S.No. | Adulterants | % (min)Limit of Detection |
| --- | --- | --- |
| 1. | Urea | 700 ppm |
| 2. | Ammonia Fertilizers | 0.1 |
| 3. | Nitrate Fertilizers /Pond Water | 0.005 |
| 4. | Starch & Cereal Flours | 0.1 |
| 5. | Sugar/Sucrose | 0.1 |
| 6. | Glucose | 0.1 |
| 7. | Salt | 0.05 |
| 8. | Neutralizers | 0.01 |
| 9. | Hydrogen Peroxide | 0.1 |
| 10. | 100 Tests Each |  |

Signature & Seal of the Tenderer

SCHEDULE – IV

PRICE BID

(To be uploaded online only)

FORM-A

Scope of Work: Supply Installation, Testing, Commissioning of Automatic Kjeldahl Unit with one year consumable parts and  3 years Annual maintenance contract (AMC) after warranty period

| Total Quantity | Rate per equipment inclusive all taxes and FOR basis (Rs in figure) | Rate per equipment inclusive all taxes and FOR basis (Rs in words) |
| --- | --- | --- |
| 11 |  |  |

FORM-B

Sope of Work: Supply Installation, Testing, Commissioning of Sodium potassium Analyses Unit with one year consumable parts and 3 years AMC  after warranty period.

| Total Quantity | Rate per equipment inclusive all taxes and FOR basis (Rs in figure) | Rate per equipments inclusive all taxes and FOR basis (Rs in words) |
| --- | --- | --- |
| 11 |  |  |

FORM-C

Scope of Work: Supply of Melamine Strips Unit.

| Total Quantity | Rate per Unit inclusive all taxes and FOR basis (Rs in figure) | Rate per Unit inclusive all taxes and FOR basis (Rs in words) |
| --- | --- | --- |
| 11 Units |  |  |

FORM-D

Scope of Work: Supply of Adulteration kit Strips Unit .

| Total Quantity | Rate per unit inclusive all taxes and FOR basis (Rs in figure) | Rate per Unit inclusive all taxes and FOR basis (Rs in words) |
| --- | --- | --- |
| 11 Units |  |  |

Signature & Seal of the Tenderer